## Slide 1
A check of ALPY calibration stability as telescope pointing changes:
German-equatorial mount
Meridian flip between target #1 and #2
Total time span of session ≈ 4 hours
Temperature wasn’t monitored. Guess drop of ≈ 10 deg F over this time span.

## Slide 2
000000
000000
N
E
1
2
3,4
5,6
7,8
meridian flip
Alt=0
Alt=60
Alt=30
Pointing to targets during night. Questions:
Are Cal amp lines stable as pointing (and temp) change?
Are Ref star Balmer lines stable?
Is Dispersion stable?

## Slide 3
Cal Lamp “red” half
5852.49Å
“red line”
Typical Cal Lamp image, showing red portion.
Measure line positions (X-coordinate) at each telescope pointing direction.

## Slide 4
Cal Lamp “blue” half
5852.49Å
“blue line”
Typical Cal Lamp image, showing blue portion

## Slide 5
peak-to-peak shift observed ≈ 0.3 pixel ≈ 1 Å
pixel coord (X)
pointing direction
Dispersion in my setup ≈ 3.35 Å/pixel

## Slide 6
ISIS ALPY routine uses both Lamp lines and Balmer lines (of Ref Star) to create Calibration polynomial:

## Slide 7
Blue-end λ-calibration using Balmer lines of a nearby reference star is better than extrapolating from a (3946Å) lamp line, but it does suffer from uncertainty in measuring line center (weak lines).
Conclusion:
ALPY is remarkably stable (<1Å) ... but do Calibrate at each new target/pointing direction.
Experiment with use of weaker Ne-AR lamp lines at blue end.

## Slide 8
broad view looks good...
observed calibrated spectrum of HD 129174 (red)
compared to MILES spectrum (blue)

## Slide 9
λ calibration at blue end looks good
observed calibrated spectrum of HD 129174 (red)
compared to MILES spectrum (blue)

## Slide 10
λ calibration at red end (H-α) matches MILES spectrum of this star.
observed calibrated spectrum of HD 129174 (red)
compared to MILES spectrum (blue)